## 总表
故障类型 | 故障类型代码 | 故障代码 | 故障代码 | 故障内容 | 故障原因 | 故障原因 | 故障位置
[7 empty] | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10
电机类故障 | 11 | 1 | MOTOR_FAIL_SETTING = 1 | 电机驱动通信失败
|  | 2 | MOTOR_UNKNOWN_FEEDBACK | 未知反馈
|  | 3 | MOTOR_RESET_FAIL | 电机驱动初始化失败
|  | 4 | MOTOR_DISABLED
|  | 5 | MOTOR_ALARM_TRIGGERED | 电机电压报警 | 1.供电不足2.电机电路或芯片损坏
|  | 6 | MOTOR_STEP_ERROR | 步进数错误 | 运行距离计算错误
|  | 7 | MOTOR_OVERCURRENT | 过电流故障 | 电机堵转、电机短路
|  | 8 | MOTOR_OVERTEMPERATURE | 过温故障
|  | 9 |  | 过热重启
编码器类故障 | 12 | 1 | ENCODER_TIMEOUT = 1 | 通信超时 | [12 empty] | 2-6上电编码值变化
|  | 2 | ENCODER_PARITY_ERROR | 偶校验失败 | [12 empty] | 2-4和2-5交换位置
|  | 3 | ENCODER_DATA_LOST | 电机运行丢步
|  | 4 | ENCODER_POWERON_FAIL | 持续无效数据 | 编码值FF,编码器损坏
|  | 5 | ENCODER_INVALID_DATA | 上电初始化失败 | 1.读存储里的编码失败
|  | 6 | ENCODER_OCF_INCOMPLETE | 上电编码值变化 | 上电编码值变化
|  | 7 | ENCODER_CORDIC_OVERFLOW | 磁性强度异常，数据无效 | 同步轮磁铁磁力过强或过弱
|  | 8 | ENCODER_LINEARITY_WARNING | 线性度报警，数据可能无效
|  | 9 |  | 相邻编码差值变化超限 | SHANG | [11 empty] | 哪一部分的数据？
|  | 10 | ENCODER_OCF_INCOMPLETE | 编码芯片角度计算未完成，数据无效 | 读取速度超过编码芯片限制1000HZ
传感器类故障 | 13 | 1 | SENSOR_BCC_ERROR = 1 | 主板数据校验失败 | 主控板到传感器校验错误，通信质量差
|  | 2 | SENSOR_COMM_TIMEOUT | 震动管频率异常 | 振动管频率扫频范围:6800-4000,如果连续三次的频率值超过6500报错
|  | 3 | SONIC_FREQ_ABNORMAL | 主板与传感器通信超时 | 无线从机到传感器
|  | 4 | DENSITY_INVALID | 密度值异常 | 0
|  | 5 | SENSOR_TEMPERATURE_ERROR | 传感器温度异常 | -200 | [11 empty] | 屏幕重启 | ·
|  | 6 | SENSOR_VOLTAGE_ERROR | 传感器电压异常
|  | 7 | 与从机通讯故障 | 无线主机与从机通讯故障
|  | 8 | 通讯过程BCC错误 | 主板与无线主机通信异常
|  | 9 | 数据包丢失率超限 | 数据包丢失率超限
|  | 10 | 无线信号强度不足 | 无线信号强度不足
测量过程类错误 | 15 | 1 | MEASUREMENT_POSITION_ERROR = 1 | 位置测量错误 | [12 empty] | 什么过程位置测量错误？
|  | 2 | MEASUREMENT_TIMEOUT | 测量超时 | [12 empty] | 是电机不懂得超市吗？
|  | 3 | MEASUREMENT_ZERO_CALIB_TIMEOUT | 标定零点超限
|  | 5 | MEASUREMENT_ZERO_REPEAT_FAIL | 零点重复性差
|  | 6 | MEASUREMENT_HEIGHT_DEVIATION | 实高偏差过大
|  | 7 | MEASUREMENT_OILLEVEL_HIGH | 液位值过高 | [12 empty] | 是到零点吗？
|  | 8 | MEASUREMENT_OILLEVEL_LOW | 向下未找到液位值
|  | 9 | MEASUREMENT_OILLEVEL_NOTFOUND | 向上未找到液位值
|  | 10 | MEASUREMENT_WEIGHT_DOWN_FAIL | 下行寻重失败 | 下行寻重失败,找罐底失败
|  | 11 | MEASUREMENT_WEIGHT_UP_FAIL | 上行寻重失败 | 上行寻重失败，找零点失败
|  | 15 | MEASUREMENT_OVERSPEED | 液位变化超速
参数/存储错误 | 17 | 1 | PARAM_EEPROM_FAIL = 1 | EEPROM写入失败
|  | 2 | PARAM_UNINITIALIZED | 参数未初始化
|  | 3 | PARAM_RANGE_ERROR | 参数范围错误
|  | 4 | PARAM_ADDRESS_OVERFLOW | 参数地址越界
|  | 5 | PARAM_CRC_ERROR | 参数读取CRC校验错误
称重故障 | 18 | 1 | WEIGHT_OUT_OF_RANGE = 1 | 称重超范围(超过上限阈值)
|  | 2 | WEIGHT_UNDER_RANGE | 称重超范围(小于下限阈值)
|  | 3 | WEIGHT_COLLISION_DETECTED | 检测到磕碰/障碍 | 50%
|  | 4 | WEIGHT_DRIFT_ERROR | 称重漂移异常 | [12 empty] | 这个的判断依据是什么？
|  | 5 | WEIGHT_SENSOR_SATURATION | 称重传感器饱和 | [12 empty] | 这个是在什么具体的居中中？
其他类型错误 | 19 | 1 | OTHER_UNKNOWN_ERROR = 1
|  | 2 | OTHER_ADDRESS_READ_ERROR
|  | 3 | OTHER_POWER_FLUCTUATION

## 设备参数
1. 基础配置参数
参数名 | 类型 | 描述 | 备注
command | CommandType | 当前指令 | 指令类型
tankHeight | uint32_t | 罐高 | 储罐总高度
blindZone | uint32_t | 盲区 | 测量盲区距离
waterBlindZone | uint32_t | 水位盲区 | 水位测量盲区
encoder_wheel_circumference_mm | uint32_t | 编码轮周长 | 编码器参数
sensorType | uint32_t | 传感器类型 | 传感器型号
softwareVersion | uint32_t | 软件版本 | 软件版本号
2. 称重参数
参数名 | 类型 | 描述 | 备注
empty_weight | uint32_t | 空载重量 | 传感器空载值
full_weight | uint32_t | 满载称重 | 传感器满载值
weight_upper_limit_ratio | uint32_t | 称重检测上限比例 | 超重保护
weight_lower_limit_ratio | uint32_t | 称重检测下限比例 | 欠重保护
empty_weight_upper_limit | uint32_t | 空载称重范围上限
empty_weight_lower_limit | uint32_t | 空载称重范围下限
full_weight_upper_limit | uint32_t | 满载称重范围上限
full_weight_lower_limit | uint32_t | 满载称重范围下限
findZeroDownDistance | uint32_t | 找零点下行距离 | 零点校准距离
3. 指令参数
参数名 | 类型 | 描述 | 备注
calibrateOilLevel | uint32_t | 标定液位值
calibrateWaterLevel | uint32_t | 水位标定值
singlePointMeasurementPosition | uint32_t | 单点测量位置
singlePointMonitoringPosition | uint32_t | 单点监测位置
densityDistributionOilLevel | uint32_t | 密度分布测量时的液位值
motorCommandDistance | uint32_t | 电机指令的运行距离
4. 密度温度测量参数
参数名 | 类型 | 描述 | 备注
densityCorrection | uint32_t | 密度修正值 | 磁通量D
temperatureCorrection | uint32_t | 温度修正值 | 磁通量T
5. 测量任务配置
参数名 | 类型 | 描述 | 备注
requireBottomMeasurement | uint32_t | 是否需要测量罐底 | 0-否,1-是
requireWaterMeasurement | uint32_t | 是否需要测量水位 | 0-否,1-是
requireSinglePointDensity | uint32_t | 是否需要测量单点密度 | 0-否,1-是
spreadMeasurementOrder | uint32_t | 分布测量顺序 | 测量顺序
spreadMeasurementMode | uint32_t | 分布测量模式 | 测量模式
spreadMeasurementCount | uint32_t | 分布测量数量 | 测量点数量
spreadMeasurementDistance | uint32_t | 分布测量间距 | 点间距
spreadTopLimit | uint32_t | 分布测量上限（距液面） | 上限距离
spreadBottomLimit | uint32_t | 分布测量下限（距罐底） | 下限距离
spreadPointHoverTime | uint32_t | 第一测量点悬停时间 | 悬停时间
6. 水位测量参数
参数名 | 类型 | 描述 | 备注
waterLevelCorrection | uint32_t | 水位修正值 | 水位校准
maxDownDistance | uint32_t | 水位测量最大下行距离 | 盲区下行距离
7. 实高测量参数
参数名 | 类型 | 描述 | 备注
refreshTankHeightFlag | uint32_t | 是否更新液位罐高 | 0-否,1-是
maxTankHeightDeviation | uint32_t | 罐高最大变化范围 | 允许偏差
initialTankHeight | uint32_t | 初始实高 | 初始罐高
currentTankHeight | uint32_t | 当前实高 | 当前罐高
8. 液位测量参数
参数名 | 类型 | 描述 | 备注
oilLevelThreshold | uint32_t | 液位跟随阈值 | 液位检测阈值
liquidLevelMeasurementMethod | uint32_t | 液位测量方式 | 测量方法
上电是否需要找液位
9. 报警输出兼容参数
参数名 | 类型 | 描述 | 备注
AlarmHighDO | uint32_t | 保留DO高兼容字段 | 当前菜单不显示，不参与继电器方式2输出
AlarmLowDO | uint32_t | 保留DO低兼容字段 | 当前菜单不显示，不参与继电器方式2输出
ThirdStateThreshold | uint32_t | 保留DO三态兼容字段 | 当前菜单不显示，不参与继电器方式2输出
10. 三路继电器方式2报警参数
参数名 | 类型 | 描述 | 备注
relayAlarm[0..2].operating_mode | uint32_t | R1~R3 工作模式 | 0=禁用；1=无源输出
relayAlarm[0..2].digital_source | uint32_t | R1~R3 数字源 | 0=无；1=高报；2=高高报；3=高/高高；4=低报；5=低低报；6=低/低低；7=全部报警
relayAlarm[0..2].contact_type | uint32_t | R1~R3 接点类型 | 0=常开；1=常闭
relayAlarm[0..2].alarm_mode | uint32_t | R1~R3 报警模式 | 0=关闭；1=开启；2=Latch
relayAlarm[0..2].error_value | uint32_t | 报警源无效时的故障报警策略 | 0=无报警；1=高高/高；2=高；3=低；4=低低/低；5=全部
relayAlarm[0..2].alarm_source | uint32_t | R1~R3 报警源 | 0=储罐液位；1=液相温度；2=水位；3=浮子位置；4=无
relayAlarm[0..2].HH_alarm_value | float | HH 报警阈值 | 32 位浮点保持寄存器；菜单小数1位
relayAlarm[0..2].H_alarm_value | float | H 报警阈值 | 32 位浮点保持寄存器；菜单小数1位
relayAlarm[0..2].L_alarm_value | float | L 报警阈值 | 32 位浮点保持寄存器；菜单小数1位
relayAlarm[0..2].LL_alarm_value | float | LL 报警阈值 | 32 位浮点保持寄存器；菜单小数1位
relayAlarm[0..2].alarm_hysteresis | float | 报警释放回差 | 用于抑制阈值附近抖动；菜单小数1位
relayAlarm[0..2].damping_factor | uint32_t | 阻尼/阻值预留参数 | 当前作为配置字段保留同步
relayAlarm[0..2].clear_alarm | uint32_t | 清除锁存报警命令 | 0=否；1=是；CPU2 消费后清回 0，不持久化
11. 4-20mA模拟输出参数
参数名 | 类型 | 描述 | 备注
CurrentRangeStart_mA | uint32_t | 电流量程起始值 | 4mA对应值
CurrentRangeEnd_mA | uint32_t | 电流量程结束值 | 20mA对应值
AlarmHighAO | uint32_t | 高液位报警输出 | 模拟输出
AlarmLowAO | uint32_t | 低液位报警输出 | 模拟输出
InitialCurrent_mA | uint32_t | 初始化电流值 | 初始电流
AOHighCurrent_mA | uint32_t | AO高报电流值 | 高报警电流
AOLowCurrent_mA | uint32_t | AO低报电流值 | 低报警电流
FaultCurrent_mA | uint32_t | 故障模式电流值 | 故障电流
DebugCurrent_mA | uint32_t | 调试模式电流值 | 调试电流
12. 校验参数
参数名 | 类型 | 描述 | 备注
crc | uint32_t | CRC校验字段 | 数据校验
瓦西莱密度梯度测量上限/下限/间距/距页面距离
调试模式
尺带拟合参数：尺带厚度，首圈周长
电机运行参数：电机运行速度，电机实际运行速度（补偿后）电机运行方向，电机运行状态
无碰撞检测上下行

## 测量结果
结构体 | 字段名 | 数据类型 | 说明
DeviceStatus | work_mode | uint32_t | 工作模式（0:工作模式 1:调试模式 78:解锁模式）
DeviceStatus | device_state | DeviceState | 设备运行状态
DeviceStatus | error_code | uint32_t | 当前错误码（0表示无错误）
DeviceStatus | current_command | CommandType | 当前指令
DeviceStatus | zero_point_status | uint8_t | 零点状态（0-正常 1-需要回零）
DensityMeasurement | temperature | uint32_t | 温度
DensityMeasurement | density | uint32_t | 密度
DensityMeasurement | temperature_position | uint32_t | 温度密度位置
DensityMeasurement | standard_density | uint32_t | 标准密度
DensityMeasurement | vcf20 | uint32_t | 体积修正系数 (VCF20)
DensityMeasurement | weight_density | uint32_t | 计重密度
DensityDistribution | average_temperature | uint32_t | 平均温度
DensityDistribution | average_density | uint32_t | 平均密度
DensityDistribution | average_standard_density | uint32_t | 平均标准密度
DensityDistribution | average_vcf20 | uint32_t | 平均体积修正系数 (VCF20)
DensityDistribution | average_weight_density | uint32_t | 平均计重密度
DensityDistribution | measurement_points | uint32_t | 实际测量点数
DensityDistribution | Density_oil_level | uint32_t | 密度分布测量时的液位值
DensityDistribution | single_density_data | DensityMeasurement[100] | 100个点的密度测量数据
ZeroCalibration | zero_encoder_turns | uint32_t | 零点编码器圈数
ZeroCalibration | zero_encoder_value | uint32_t | 零点编码器编码值
DebugData | current_encoder_value | int32_t | 当前编码值
DebugData | sensor_position | int32_t | 传感器位置
DebugData | cable_length | int32_t | 尺带长度
DebugData | frequency | uint32_t | 频率
DebugData | temperature | uint32_t | 温度 (单位:0.01°C)
DebugData | air_frequency | uint32_t | 空气中频率
DebugData | current_amplitude | uint32_t | 当前幅值
DebugData | water_level_voltage | uint32_t | 水位电压值/电容值
当前称重值
X轴角度值
Y轴角度值
称重参数
ActualHeightMeasurement | calibrated_liquid_level | uint32_t | 标定液位时实高
ActualHeightMeasurement | current_real_height | uint32_t | 当前实高
OilMeasurement | oil_level | uint32_t | 液位跟随液位值
OilMeasurement | air_frequency | uint32_t | 空气中频率
OilMeasurement | oil_frequency | uint32_t | 油中频率
OilMeasurement | follow_frequency | uint32_t | 液位跟随频率
OilMeasurement | current_frequency | uint32_t | 当前频率
WaterMeasurement | water_level | uint32_t | 测量水位的值
WaterMeasurement | zero_capacitance | float | 零点电容
WaterMeasurement | oil_capacitance | float | 油中电容
WaterMeasurement | current_capacitance | float | 当前电容
MeasurementResult | device_status | DeviceStatus | 设备状态
MeasurementResult | oil_measurement | OilMeasurement | 液位测量数据
MeasurementResult | water_measurement | WaterMeasurement | 水位测量数据
MeasurementResult | zero_calibration | ZeroCalibration | 标定零点数据
MeasurementResult | height_measurement | ActualHeightMeasurement | 实高测量数据
MeasurementResult | debug_data | DebugData | 调试数据
MeasurementResult | single_point_measurement | DensityMeasurement | 单点测量数据
MeasurementResult | single_point_monitoring | DensityMeasurement | 单点监测数据
MeasurementResult | density_distribution | DensityDistribution | 密度分布测量数据

## 保持寄存器表显示 (hidden)
struct HoldRegisterData holdValue[] = {
| /* 名称 | 默认 | 数据号 | 起始地址 | 寄存器数 | 是否检 | 最小 | 最大 | 单位 | 小数 | 偏移 | 写权 | 类型 | 显示 | 隐藏 | 英文 */
| {(uint8_t*)"设备指令", | 0, | COM_NUM_DEVICEPARAM_COMMAND, | HOLDREGISTER_DEVICEPARAM_COMMAND, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Cmd"},
/* 基础参数 */
| {(uint8_t*)"罐高", | 0, | COM_NUM_DEVICEPARAM_TANKHEIGHT, | HOLDREGISTER_DEVICEPARAM_TANKHEIGHT, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"TankHeight"},
| {(uint8_t*)"液位盲区", | 0, | COM_NUM_DEVICEPARAM_BLINDZONE, | HOLDREGISTER_DEVICEPARAM_BLINDZONE, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"BlindZone"},
| {(uint8_t*)"水位盲区", | 0, | COM_NUM_DEVICEPARAM_WATER_BLINDZONE, | HOLDREGISTER_DEVICEPARAM_WATER_BLINDZONE, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"WaterBlind"},
| {(uint8_t*)"编码轮周长mm", | 0, | COM_NUM_DEVICEPARAM_ENCODER_WHEEL_CIRCUMFERENCE_MM, | HOLDREGISTER_DEVICEPARAM_ENCODER_WHEEL_CIRCUMFERENCE_MM, | 2, | false, | 0, | 0, | NULL, | 3, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"EncWheelCirc"},
| {(uint8_t*)"传感器类型", | 0, | COM_NUM_DEVICEPARAM_SENSORTYPE, | HOLDREGISTER_DEVICEPARAM_SENSORTYPE, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 3, | NULL, | (uint8_t*)"SensorType"},
| {(uint8_t*)"传感器编号", | 0, | COM_NUM_DEVICEPARAM_SENSORID, | HOLDREGISTER_DEVICEPARAM_SENSORID, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"SensorID"},
| {(uint8_t*)"软件版本", | 0, | COM_NUM_DEVICEPARAM_SOFTWAREVERSION, | HOLDREGISTER_DEVICEPARAM_SOFTWAREVERSION, | 2, | false, | 0, | 0, | NULL, | 3, | 0, | false, | TYPE_INT, | 7, | NULL, | (uint8_t*)"FWVersion"},
/* 称重参数 */
| {(uint8_t*)"空桶重量", | 0, | COM_NUM_DEVICEPARAM_EMPTY_WEIGHT, | HOLDREGISTER_DEVICEPARAM_EMPTY_WEIGHT, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"EmptyWt"},
| {(uint8_t*)"满桶重量", | 0, | COM_NUM_DEVICEPARAM_FULL_WEIGHT, | HOLDREGISTER_DEVICEPARAM_FULL_WEIGHT, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"FullWt"},
| {(uint8_t*)"上限比率", | 0, | COM_NUM_DEVICEPARAM_WEIGHT_UPPER_LIMIT_RATIO, | HOLDREGISTER_DEVICEPARAM_WEIGHT_UPPER_LIMIT_RATIO, | 2, | false, | 0, | 0, | (uint8_t*)"%", | 0, | 0, | true, | TYPE_INT, | 3, | NULL, | (uint8_t*)"UpperRatio"},
| {(uint8_t*)"下限比率", | 0, | COM_NUM_DEVICEPARAM_WEIGHT_LOWER_LIMIT_RATIO, | HOLDREGISTER_DEVICEPARAM_WEIGHT_LOWER_LIMIT_RATIO, | 2, | false, | 0, | 0, | (uint8_t*)"%", | 0, | 0, | true, | TYPE_INT, | 3, | NULL, | (uint8_t*)"LowerRatio"},
| {(uint8_t*)"空桶重量上限", | 0, | COM_NUM_DEVICEPARAM_EMPTY_WEIGHT_UPPER_LIMIT, | HOLDREGISTER_DEVICEPARAM_EMPTY_WEIGHT_UPPER_LIMIT, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"EmptyWtHi"},
| {(uint8_t*)"空桶重量下限", | 0, | COM_NUM_DEVICEPARAM_EMPTY_WEIGHT_LOWER_LIMIT, | HOLDREGISTER_DEVICEPARAM_EMPTY_WEIGHT_LOWER_LIMIT, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"EmptyWtLo"},
| {(uint8_t*)"满桶重量上限", | 0, | COM_NUM_DEVICEPARAM_FULL_WEIGHT_UPPER_LIMIT, | HOLDREGISTER_DEVICEPARAM_FULL_WEIGHT_UPPER_LIMIT, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"FullWtHi"},
| {(uint8_t*)"满桶重量下限", | 0, | COM_NUM_DEVICEPARAM_FULL_WEIGHT_LOWER_LIMIT, | HOLDREGISTER_DEVICEPARAM_FULL_WEIGHT_LOWER_LIMIT, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"FullWtLo"},
| {(uint8_t*)"找零下行距离", | 0, | COM_NUM_DEVICEPARAM_FINDZERO_DOWN_DISTANCE, | HOLDREGISTER_DEVICEPARAM_FINDZERO_DOWN_DISTANCE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"FindZeroDown"},
/* 指令参数 */
| {(uint8_t*)"标定液位(油)", | 0, | COM_NUM_CAL_OIL, | HOLDREGISTER_DEVICEPARAM_CALIBRATE_OIL_LEVEL, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"CalOilLvl"},
| {(uint8_t*)"修正液位(油)", | 0, | COM_NUM_CORRECTION_OIL, | HOLDREGISTER_DEVICEPARAM_CALIBRATE_OIL_LEVEL, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"CorrectLvl"},
| {(uint8_t*)"标定液位(水)", | 0, | COM_NUM_DEVICEPARAM_CALIBRATE_WATER_LEVEL, | HOLDREGISTER_DEVICEPARAM_CALIBRATE_WATER_LEVEL, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"CalWaterLvl"},
| {(uint8_t*)"单点测量位置", | 0, | COM_NUM_SINGLE_POINT, | HOLDREGISTER_DEVICEPARAM_SP_MEAS_POSITION, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"SP_MeasPos"},
| {(uint8_t*)"单点监测位置", | 0, | COM_NUM_SP_TEST, | HOLDREGISTER_DEVICEPARAM_SP_MONITOR_POSITION, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"SP_MonPos"},
| {(uint8_t*)"分布测液位", | 0, | COM_NUM_SPREADPOINTS, | HOLDREGISTER_DEVICEPARAM_DENSITY_DISTRIBUTION_OIL_LEVEL, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"DistLvl"},
| {(uint8_t*)"电机上行距离", | 0, | COM_NUM_RUNUP, | HOLDREGISTER_DEVICEPARAM_MOTOR_COMMAND_DISTANCE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"MotorUp"},
| {(uint8_t*)"电机下行距离", | 0, | COM_NUM_RUNDOWN, | HOLDREGISTER_DEVICEPARAM_MOTOR_COMMAND_DISTANCE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"MotorDown"},
/* 修正参数 */
| {(uint8_t*)"密度修正", | 0, | COM_NUM_DEVICEPARAM_DENSITYCORRECTION, | HOLDREGISTER_DEVICEPARAM_DENSITYCORRECTION, | 2, | false, | -10000, | 10000, | NULL, | 1, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"DensityCorr"},
| {(uint8_t*)"温度修正", | 0, | COM_NUM_DEVICEPARAM_TEMPERATURECORRECTION, | HOLDREGISTER_DEVICEPARAM_TEMPERATURECORRECTION, | 2, | false, | -1000, | 1000, | NULL, | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"TempCorr"},
/* 分布测量参数 */
| {(uint8_t*)"是否测罐底", | 0, | COM_NUM_DEVICEPARAM_REQUIREBOTTOMMEASUREMENT, | HOLDREGISTER_DEVICEPARAM_REQUIREBOTTOMMEASUREMENT, | 2, | false, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"NeedBottom"},
| {(uint8_t*)"是否测水位", | 0, | COM_NUM_DEVICEPARAM_REQUIREWATERMEASUREMENT, | HOLDREGISTER_DEVICEPARAM_REQUIREWATERMEASUREMENT, | 2, | false, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"NeedWater"},
| {(uint8_t*)"是否测单点密度", | 0, | COM_NUM_DEVICEPARAM_REQUIRESINGLEPOINTDENSITY, | HOLDREGISTER_DEVICEPARAM_REQUIRESINGLEPOINTDENSITY, | 2, | false, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"NeedSingleDens"},
| {(uint8_t*)"分布测顺序", | 0, | COM_NUM_DEVICEPARAM_SPREADMEASUREMENTORDER, | HOLDREGISTER_DEVICEPARAM_SPREADMEASUREMENTORDER, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"SpreadOrder"},
| {(uint8_t*)"分布测模式", | 0, | COM_NUM_DEVICEPARAM_SPREADMEASUREMENTMODE, | HOLDREGISTER_DEVICEPARAM_SPREADMEASUREMENTMODE, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"SpreadMode"},
| {(uint8_t*)"分布测点数", | 0, | COM_NUM_DEVICEPARAM_SPREADMEASUREMENTCOUNT, | HOLDREGISTER_DEVICEPARAM_SPREADMEASUREMENTCOUNT, | 2, | false, | 1, | 99, | NULL, | 0, | 0, | true, | TYPE_INT, | 3, | NULL, | (uint8_t*)"SpreadCount"},
| {(uint8_t*)"分布点间距", | 0, | COM_NUM_DEVICEPARAM_SPREADMEASUREMENTDISTANCE, | HOLDREGISTER_DEVICEPARAM_SPREADMEASUREMENTDISTANCE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"SpreadDist"},
| {(uint8_t*)"顶点距液面", | 0, | COM_NUM_DEVICEPARAM_SPREADTOPLIMIT, | HOLDREGISTER_DEVICEPARAM_SPREADTOPLIMIT, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"Top2Level"},
| {(uint8_t*)"底点距罐底", | 0, | COM_NUM_DEVICEPARAM_SPREADBOTTOMLIMIT, | HOLDREGISTER_DEVICEPARAM_SPREADBOTTOMLIMIT, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"Bot2Tank"},
| {(uint8_t*)"分布点悬停(s)", | 0, | COM_NUM_DEVICEPARAM_SPREAD_POINT_HOVER_TIME, | HOLDREGISTER_DEVICEPARAM_SPREAD_POINT_HOVER_TIME, | 2, | false, | 0, | 600, | (uint8_t*)"s", | 0, | 0, | true, | TYPE_INT, | 3, | NULL, | (uint8_t*)"PtHoverTime"},
/* 水位测量 */
| {(uint8_t*)"水位修正值", | 0, | COM_NUM_DEVICEPARAM_WATERLEVELCORRECTION, | HOLDREGISTER_DEVICEPARAM_WATERLEVELCORRECTION, | 2, | false, | -1000, | 1000, | NULL, | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"WaterCorr"},
| {(uint8_t*)"最大下行距离", | 0, | COM_NUM_DEVICEPARAM_MAXDOWNDISTANCE, | HOLDREGISTER_DEVICEPARAM_MAXDOWNDISTANCE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"MaxDown"},
/* 实高测量 */
| {(uint8_t*)"更新罐高标志", | 0, | COM_NUM_DEVICEPARAM_REFRESHTANKHEIGHTFLAG, | HOLDREGISTER_DEVICEPARAM_REFRESHTANKHEIGHTFLAG, | 2, | false, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"UpdTankHeight"},
| {(uint8_t*)"实高最大偏差", | 0, | COM_NUM_DEVICEPARAM_MAXTANKHEIGHTDEVIATION, | HOLDREGISTER_DEVICEPARAM_MAXTANKHEIGHTDEVIATION, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"MaxDev"},
| {(uint8_t*)"初始罐高", | 0, | COM_NUM_DEVICEPARAM_INITIALTANKHEIGHT, | HOLDREGISTER_DEVICEPARAM_INITIALTANKHEIGHT, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | false, | TYPE_INT, | 7, | NULL, | (uint8_t*)"InitTankH"},
| {(uint8_t*)"当前罐高", | 0, | COM_NUM_DEVICEPARAM_CURRENTTANKHEIGHT, | HOLDREGISTER_DEVICEPARAM_CURRENTTANKHEIGHT, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | false, | TYPE_INT, | 7, | NULL, | (uint8_t*)"CurrTankH"},
/* 液位测量 */
| {(uint8_t*)"找油阈值", | 0, | COM_NUM_DEVICEPARAM_OILLEVELTHRESHOLD, | HOLDREGISTER_DEVICEPARAM_OILLEVELTHRESHOLD, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"OilLvlTh"},
| {(uint8_t*)"液位测量方式", | 0, | COM_NUM_DEVICEPARAM_LIQUIDLEVELMEASUREMENTMETHOD, | HOLDREGISTER_DEVICEPARAM_LIQUIDLEVELMEASUREMENTMETHOD, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"LvlMeasMode"},
/* 报警（DO） */
| {(uint8_t*)"保留DO高", | 0, | COM_NUM_DEVICEPARAM_ALARM_HIGH_DO, | HOLDREGISTER_DEVICEPARAM_ALARM_HIGH_DO, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"AlarmHiDO"},
| {(uint8_t*)"保留DO低", | 0, | COM_NUM_DEVICEPARAM_ALARM_LOW_DO, | HOLDREGISTER_DEVICEPARAM_ALARM_LOW_DO, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"AlarmLoDO"},
| {(uint8_t*)"保留DO三态", | 0, | COM_NUM_DEVICEPARAM_THIRD_STATE_THRESHOLD, | HOLDREGISTER_DEVICEPARAM_THIRD_STATE_THRESHOLD, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"ThirdStateTh"},
/* 4–20mA 输出与报警（AO） */
| {(uint8_t*)"输出范围起点mA", | 0, | COM_NUM_DEVICEPARAM_CURRENT_RANGE_START_mA, | HOLDREGISTER_DEVICEPARAM_CURRENT_RANGE_START_mA, | 2, | false, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"RangeStart"},
| {(uint8_t*)"输出范围终点mA", | 0, | COM_NUM_DEVICEPARAM_CURRENT_RANGE_END_mA, | HOLDREGISTER_DEVICEPARAM_CURRENT_RANGE_END_mA, | 2, | false, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"RangeEnd"},
| {(uint8_t*)"高限报警(AO)", | 0, | COM_NUM_DEVICEPARAM_ALARM_HIGH_AO, | HOLDREGISTER_DEVICEPARAM_ALARM_HIGH_AO, | 2, | false, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"AlarmHiAO"},
| {(uint8_t*)"低限报警(AO)", | 0, | COM_NUM_DEVICEPARAM_ALARM_LOW_AO, | HOLDREGISTER_DEVICEPARAM_ALARM_LOW_AO, | 2, | false, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"AlarmLoAO"},
| {(uint8_t*)"初始电流mA", | 0, | COM_NUM_DEVICEPARAM_INITIAL_CURRENT_mA, | HOLDREGISTER_DEVICEPARAM_INITIAL_CURRENT_mA, | 2, | false, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"InitCurrent"},
| {(uint8_t*)"高位电流mA", | 0, | COM_NUM_DEVICEPARAM_AO_HIGH_CURRENT_mA, | HOLDREGISTER_DEVICEPARAM_AO_HIGH_CURRENT_mA, | 2, | false, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"HighCurrent"},
| {(uint8_t*)"低位电流mA", | 0, | COM_NUM_DEVICEPARAM_AO_LOW_CURRENT_mA, | HOLDREGISTER_DEVICEPARAM_AO_LOW_CURRENT_mA, | 2, | false, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"LowCurrent"},
| {(uint8_t*)"故障电流mA", | 0, | COM_NUM_DEVICEPARAM_FAULT_CURRENT_mA, | HOLDREGISTER_DEVICEPARAM_FAULT_CURRENT_mA, | 2, | false, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"FaultCurrent"},
| {(uint8_t*)"调试电流mA", | 0, | COM_NUM_DEVICEPARAM_DEBUG_CURRENT_mA, | HOLDREGISTER_DEVICEPARAM_DEBUG_CURRENT_mA, | 2, | false, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"DebugCurrent"},
/* CRC */
| {(uint8_t*)"参数CRC32", | 0, | COM_NUM_DEVICEPARAM_CRC, | HOLDREGISTER_DEVICEPARAM_CRC, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"ParamCRC"},
};

## 保持寄存器显示
| /* 名称 | 数值 | 数据号 | 起始地址 | 寄存器数 | 是否检范围 | 最小 | 最大 | 单位 | 小数 | 偏移 | 写权 | 类型 | 显示 | 隐藏 | 英文 */
//参数元数据
struct ParameterMetadata param_meta[] = {
| {(uint8_t*)"设备指令", | 0, | COM_NUM_DEVICEPARAM_COMMAND, | HOLDREGISTER_DEVICEPARAM_COMMAND, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"Cmd"},
| {(uint8_t*)"传感器类型", | 0, | COM_NUM_DEVICEPARAM_SENSORTYPE, | HOLDREGISTER_DEVICEPARAM_SENSORTYPE, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 3, | NULL, | (uint8_t*)"SensorType"},
| {(uint8_t*)"传感器编号", | 0, | COM_NUM_DEVICEPARAM_SENSORID, | HOLDREGISTER_DEVICEPARAM_SENSORID, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"SensorID"},
| {(uint8_t*)"传感器软件版本", | 0, | COM_NUM_DEVICEPARAM_SENSOR_SOFTWARE_VERSION, | HOLDREGISTER_DEVICEPARAM_SENSOR_SOFTWARE_VERSION, | 2, | false, | 0, | 0, | NULL, | 3, | 0, | false, | TYPE_INT, | 4, | NULL, | (uint8_t*)"SenSWVer"},
| {(uint8_t*)"软件版本", | 0, | COM_NUM_DEVICEPARAM_SOFTWAREVERSION, | HOLDREGISTER_DEVICEPARAM_SOFTWAREVERSION, | 2, | false, | 0, | 0, | NULL, | 3, | 0, | false, | TYPE_INT, | 7, | NULL, | (uint8_t*)"FWVersion"},
| {(uint8_t*)"上电默认指令", | 0, | COM_NUM_DEVICEPARAM_POWER_ON_DEFAULT_COMMAND, | HOLDREGISTER_DEVICEPARAM_POWER_ON_DEFAULT_COMMAND, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 3, | ret_arr_word, | (uint8_t*)"PwrOnCmd"},
| {(uint8_t*)"故障自动回零", | 0, | COM_NUM_DEVICEPARAM_ERROR_AUTO_BACK_ZERO, | HOLDREGISTER_DEVICEPARAM_ERROR_AUTO_BACK_ZERO, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"ErrAutoBackZero"},
| {(uint8_t*)"故障停止测量", | 0, | COM_NUM_DEVICEPARAM_ERROR_STOP_MEASUREMENT, | HOLDREGISTER_DEVICEPARAM_ERROR_STOP_MEASUREMENT, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"ErrStopMeas"},
| {(uint8_t*)"保留1", | 0, | COM_NUM_DEVICEPARAM_RESERVED1, | HOLDREGISTER_DEVICEPARAM_RESERVED1, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv1"},
| {(uint8_t*)"保留2", | 0, | COM_NUM_DEVICEPARAM_RESERVED2, | HOLDREGISTER_DEVICEPARAM_RESERVED2, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv2"},
| {(uint8_t*)"保留3", | 0, | COM_NUM_DEVICEPARAM_RESERVED3, | HOLDREGISTER_DEVICEPARAM_RESERVED3, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv3"},
| {(uint8_t*)"保留4", | 0, | COM_NUM_DEVICEPARAM_RESERVED4, | HOLDREGISTER_DEVICEPARAM_RESERVED4, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv4"},
| {(uint8_t*)"保留5", | 0, | COM_NUM_DEVICEPARAM_RESERVED5, | HOLDREGISTER_DEVICEPARAM_RESERVED5, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv5"},
| {(uint8_t*)"编码轮周长mm", | 0, | COM_NUM_DEVICEPARAM_ENCODER_WHEEL_CIRCUMFERENCE_MM, | HOLDREGISTER_DEVICEPARAM_ENCODER_WHEEL_CIRCUMFERENCE_MM, | 2, | false, | 0, | 0, | NULL, | 3, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"EncWheelCirc"},
| {(uint8_t*)"电机最大速度", | 0, | COM_NUM_DEVICEPARAM_MAX_MOTOR_SPEED, | HOLDREGISTER_DEVICEPARAM_MAX_MOTOR_SPEED, | 2, | false, | 0, | 0, | NULL, | 2, | 0, | true, | TYPE_INT, | 3, | NULL, | (uint8_t*)"MaxMotorSpd"},
| {(uint8_t*)"首圈周长mm", | 0, | COM_NUM_DEVICEPARAM_FIRST_LOOP_CIRCUMFERENCE_MM, | HOLDREGISTER_DEVICEPARAM_FIRST_LOOP_CIRCUMFERENCE_MM, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"FirstLoop"},
| {(uint8_t*)"尺带厚度mm", | 0, | COM_NUM_DEVICEPARAM_TAPE_THICKNESS_MM, | HOLDREGISTER_DEVICEPARAM_TAPE_THICKNESS_MM, | 2, | false, | 0, | 0, | NULL, | 3, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"TapeThick"},
| {(uint8_t*)"保留6", | 0, | COM_NUM_DEVICEPARAM_RESERVED6, | HOLDREGISTER_DEVICEPARAM_RESERVED6, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv6"},
| {(uint8_t*)"保留7", | 0, | COM_NUM_DEVICEPARAM_RESERVED7, | HOLDREGISTER_DEVICEPARAM_RESERVED7, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv7"},
| {(uint8_t*)"空载重量", | 0, | COM_NUM_DEVICEPARAM_EMPTY_WEIGHT, | HOLDREGISTER_DEVICEPARAM_EMPTY_WEIGHT, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"EmptyWt"},
| {(uint8_t*)"空载重量上限", | 0, | COM_NUM_DEVICEPARAM_EMPTY_WEIGHT_UPPER_LIMIT, | HOLDREGISTER_DEVICEPARAM_EMPTY_WEIGHT_UPPER_LIMIT, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"EmptyWtHi"},
| {(uint8_t*)"空载重量下限", | 0, | COM_NUM_DEVICEPARAM_EMPTY_WEIGHT_LOWER_LIMIT, | HOLDREGISTER_DEVICEPARAM_EMPTY_WEIGHT_LOWER_LIMIT, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"EmptyWtLo"},
| {(uint8_t*)"满载重量", | 0, | COM_NUM_DEVICEPARAM_FULL_WEIGHT, | HOLDREGISTER_DEVICEPARAM_FULL_WEIGHT, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"FullWt"},
| {(uint8_t*)"满载重量上限", | 0, | COM_NUM_DEVICEPARAM_FULL_WEIGHT_UPPER_LIMIT, | HOLDREGISTER_DEVICEPARAM_FULL_WEIGHT_UPPER_LIMIT, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"FullWtHi"},
| {(uint8_t*)"满载重量下限", | 0, | COM_NUM_DEVICEPARAM_FULL_WEIGHT_LOWER_LIMIT, | HOLDREGISTER_DEVICEPARAM_FULL_WEIGHT_LOWER_LIMIT, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"FullWtLo"},
| {(uint8_t*)"碰撞上限比率", | 0, | COM_NUM_DEVICEPARAM_WEIGHT_UPPER_LIMIT_RATIO, | HOLDREGISTER_DEVICEPARAM_WEIGHT_UPPER_LIMIT_RATIO, | 2, | false, | 0, | 0, | (uint8_t*)"%", | 0, | 0, | true, | TYPE_INT, | 3, | NULL, | (uint8_t*)"UpperRatio"},
| {(uint8_t*)"碰撞下限比率", | 0, | COM_NUM_DEVICEPARAM_WEIGHT_LOWER_LIMIT_RATIO, | HOLDREGISTER_DEVICEPARAM_WEIGHT_LOWER_LIMIT_RATIO, | 2, | false, | 0, | 0, | (uint8_t*)"%", | 0, | 0, | true, | TYPE_INT, | 3, | NULL, | (uint8_t*)"LowerRatio"},
| {(uint8_t*)"保留8", | 0, | COM_NUM_DEVICEPARAM_RESERVED8, | HOLDREGISTER_DEVICEPARAM_RESERVED8, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv8"},
| {(uint8_t*)"保留9", | 0, | COM_NUM_DEVICEPARAM_RESERVED9, | HOLDREGISTER_DEVICEPARAM_RESERVED9, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv9"},
| {(uint8_t*)"零点阈值比例", | 0, | COM_NUM_DEVICEPARAM_ZERO_WEIGHT_THRESHOLD_RATIO, | HOLDREGISTER_DEVICEPARAM_ZERO_WEIGHT_THRESHOLD_RATIO, | 2, | true, | 0, | 100, | (uint8_t*)"%", | 0, | 0, | true, | TYPE_INT, | 3, | NULL, | (uint8_t*)"ZeroThRatio"},
| {(uint8_t*)"称重忽略区", | 0, | COM_NUM_DEVICEPARAM_WEIGHT_IGNORE_ZONE, | HOLDREGISTER_DEVICEPARAM_WEIGHT_IGNORE_ZONE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"WeightIgnore"},
| {(uint8_t*)"零点最大偏差", | 0, | COM_NUM_DEVICEPARAM_MAX_ZERO_DEVIATION_DISTANCE, | HOLDREGISTER_DEVICEPARAM_MAX_ZERO_DEVIATION_DISTANCE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"MaxZeroDev"},
| {(uint8_t*)"找零下行距离", | 0, | COM_NUM_DEVICEPARAM_FINDZERO_DOWN_DISTANCE, | HOLDREGISTER_DEVICEPARAM_FINDZERO_DOWN_DISTANCE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"FindZeroDown"},
| {(uint8_t*)"保留10", | 0, | COM_NUM_DEVICEPARAM_RESERVED10, | HOLDREGISTER_DEVICEPARAM_RESERVED10, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv10"},
| {(uint8_t*)"保留11", | 0, | COM_NUM_DEVICEPARAM_RESERVED11, | HOLDREGISTER_DEVICEPARAM_RESERVED11, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv11"},
| {(uint8_t*)"液位罐高", | 0, | COM_NUM_DEVICEPARAM_TANKHEIGHT, | HOLDREGISTER_DEVICEPARAM_TANKHEIGHT, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"TankHeight"},
| {(uint8_t*)"液位探头距差", | 0, | COM_NUM_DEVICEPARAM_LIQUID_SENSOR_DISTANCE_DIFF, | HOLDREGISTER_DEVICEPARAM_LIQUID_SENSOR_DISTANCE_DIFF, | 2, | false, | -999999, | 999999, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"LvlSenDiff"},
| {(uint8_t*)"液位盲区", | 0, | COM_NUM_DEVICEPARAM_BLINDZONE, | HOLDREGISTER_DEVICEPARAM_BLINDZONE, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"BlindZone"},
| {(uint8_t*)"找油阈值", | 0, | COM_NUM_DEVICEPARAM_OILLEVELTHRESHOLD, | HOLDREGISTER_DEVICEPARAM_OILLEVELTHRESHOLD, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"OilLvlTh"},
| {(uint8_t*)"滞后阈值", | 0, | COM_NUM_DEVICEPARAM_OILLEVEL_HYSTERESIS_THRESHOLD, | HOLDREGISTER_DEVICEPARAM_OILLEVEL_HYSTERESIS_THRESHOLD, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"HysTh"},
| {(uint8_t*)"液位测量方式", | 0, | COM_NUM_DEVICEPARAM_LIQUIDLEVELMEASUREMENTMETHOD, | HOLDREGISTER_DEVICEPARAM_LIQUIDLEVELMEASUREMENTMETHOD, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"LvlMeasMode"},
| {(uint8_t*)"液位跟随频率", | 0, | COM_NUM_DEVICEPARAM_OILLEVEL_FREQUENCY, | HOLDREGISTER_DEVICEPARAM_OILLEVEL_FREQUENCY, | 2, | false, | 0, | 0, | (uint8_t*)"HZ", | 0, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"FollowFreq"},
| {(uint8_t*)"液位跟随密度", | 0, | COM_NUM_DEVICEPARAM_OILLEVEL_DENSITY, | HOLDREGISTER_DEVICEPARAM_OILLEVEL_DENSITY, | 2, | false, | 0, | 0, | (uint8_t*)"kg/m3", | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"FollowDens"},
| {(uint8_t*)"水位罐高", | 0, | COM_NUM_DEVICEPARAM_WATER_TANK_HEIGHT, | HOLDREGISTER_DEVICEPARAM_WATER_TANK_HEIGHT, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"WaterTankH"},
| {(uint8_t*)"水位测量方式", | 0, | COM_NUM_DEVICEPARAM_WATER_LEVEL_MODE, | HOLDREGISTER_DEVICEPARAM_WATER_LEVEL_MODE, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"WaterSenDiff"},
| {(uint8_t*)"水位盲区", | 0, | COM_NUM_DEVICEPARAM_WATER_BLINDZONE, | HOLDREGISTER_DEVICEPARAM_WATER_BLINDZONE, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"WaterBlind"},
| {(uint8_t*)"水位电容阈值", | 0, | COM_NUM_DEVICEPARAM_WATER_CAP_THRESHOLD, | HOLDREGISTER_DEVICEPARAM_WATER_CAP_THRESHOLD, | 2, | false, | 0, | 0, | NULL, | 3, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"WaterCapTh"},
| {(uint8_t*)"水位电容滞回", | 0, | COM_NUM_DEVICEPARAM_WATER_CAP_HYSTERESIS, | HOLDREGISTER_DEVICEPARAM_WATER_CAP_HYSTERESIS, | 2, | false, | 0, | 0, | NULL, | 3, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"WaterCapHys"},
| {(uint8_t*)"最大下行距离", | 0, | COM_NUM_DEVICEPARAM_MAXDOWNDISTANCE, | HOLDREGISTER_DEVICEPARAM_MAXDOWNDISTANCE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"MaxDown"},
| {(uint8_t*)"零点电容", | 0, | COM_NUM_DEVICEPARAM_ZERO_CAP, | HOLDREGISTER_DEVICEPARAM_ZERO_CAP, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"ZeroCap"},
| {(uint8_t*)"水位稳定阈值", | 0, | COM_NUM_DEVICEPARAM_WATER_STABLE_THRESHOLD, | HOLDREGISTER_DEVICEPARAM_WATER_STABLE_THRESHOLD, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"Rsv15"},
| {(uint8_t*)"罐底检测模式", | 0, | COM_NUM_DEVICEPARAM_BOTTOM_DETECT_MODE, | HOLDREGISTER_DEVICEPARAM_BOTTOM_DETECT_MODE, | 2, | true, | 0, | 3, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"BottomMode"},
| {(uint8_t*)"罐底角度阈值", | 0, | COM_NUM_DEVICEPARAM_BOTTOM_ANGLE_THRESHOLD, | HOLDREGISTER_DEVICEPARAM_BOTTOM_ANGLE_THRESHOLD, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"BottomAngTh"},
| {(uint8_t*)"罐底称重阈值", | 0, | COM_NUM_DEVICEPARAM_BOTTOM_WEIGHT_THRESHOLD, | HOLDREGISTER_DEVICEPARAM_BOTTOM_WEIGHT_THRESHOLD, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"BottomWtTh"},
| {(uint8_t*)"更新罐高标志", | 0, | COM_NUM_DEVICEPARAM_REFRESH_TANKHEIGHT_FLAG, | HOLDREGISTER_DEVICEPARAM_REFRESH_TANKHEIGHT_FLAG, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"UpdTankHeight"},
| {(uint8_t*)"实高最大偏差", | 0, | COM_NUM_DEVICEPARAM_MAX_TANKHEIGHT_DEVIATION, | HOLDREGISTER_DEVICEPARAM_MAX_TANKHEIGHT_DEVIATION, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"MaxDev"},
| {(uint8_t*)"初始罐高", | 0, | COM_NUM_DEVICEPARAM_INITIAL_TANKHEIGHT, | HOLDREGISTER_DEVICEPARAM_INITIAL_TANKHEIGHT, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | false, | TYPE_INT, | 7, | NULL, | (uint8_t*)"InitTankH"},
| {(uint8_t*)"当前罐高", | 0, | COM_NUM_DEVICEPARAM_CURRENT_TANKHEIGHT, | HOLDREGISTER_DEVICEPARAM_CURRENT_TANKHEIGHT, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | false, | TYPE_INT, | 7, | NULL, | (uint8_t*)"CurrTankH"},
| {(uint8_t*)"保留16", | 0, | COM_NUM_DEVICEPARAM_RESERVED16, | HOLDREGISTER_DEVICEPARAM_RESERVED16, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv16"},
| {(uint8_t*)"保留17", | 0, | COM_NUM_DEVICEPARAM_RESERVED17, | HOLDREGISTER_DEVICEPARAM_RESERVED17, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv17"},
| {(uint8_t*)"磁通量D", | 0, | COM_NUM_DEVICEPARAM_DENSITYCORRECTION, | HOLDREGISTER_DEVICEPARAM_DENSITYCORRECTION, | 2, | false, | -10000, | 10000, | NULL, | 1, | -10000, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"DensityCorr"},
| {(uint8_t*)"磁通量T", | 0, | COM_NUM_DEVICEPARAM_TEMPERATURECORRECTION, | HOLDREGISTER_DEVICEPARAM_TEMPERATURECORRECTION, | 2, | false, | -1000, | 1000, | NULL, | 1, | -1000, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"TempCorr"},
| {(uint8_t*)"保留18", | 0, | COM_NUM_DEVICEPARAM_RESERVED18, | HOLDREGISTER_DEVICEPARAM_RESERVED18, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv18"},
| {(uint8_t*)"保留19", | 0, | COM_NUM_DEVICEPARAM_RESERVED19, | HOLDREGISTER_DEVICEPARAM_RESERVED19, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv19"},
| {(uint8_t*)"是否测罐底", | 0, | COM_NUM_DEVICEPARAM_REQUIREBOTTOMMEASUREMENT, | HOLDREGISTER_DEVICEPARAM_REQUIREBOTTOMMEASUREMENT, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"NeedBottom"},
| {(uint8_t*)"是否测水位", | 0, | COM_NUM_DEVICEPARAM_REQUIREWATERMEASUREMENT, | HOLDREGISTER_DEVICEPARAM_REQUIREWATERMEASUREMENT, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"NeedWater"},
| {(uint8_t*)"是否测单点密度", | 0, | COM_NUM_DEVICEPARAM_REQUIRESINGLEPOINTDENSITY, | HOLDREGISTER_DEVICEPARAM_REQUIRESINGLEPOINTDENSITY, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"NeedSingleDens"},
| {(uint8_t*)"分布测顺序", | 0, | COM_NUM_DEVICEPARAM_SPREADMEASUREMENTORDER, | HOLDREGISTER_DEVICEPARAM_SPREADMEASUREMENTORDER, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"SpreadOrder"},
| {(uint8_t*)"分布测模式", | 0, | COM_NUM_DEVICEPARAM_SPREADMEASUREMENTMODE, | HOLDREGISTER_DEVICEPARAM_SPREADMEASUREMENTMODE, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"SpreadMode"},
| {(uint8_t*)"分布测点数", | 0, | COM_NUM_DEVICEPARAM_SPREADMEASUREMENTCOUNT, | HOLDREGISTER_DEVICEPARAM_SPREADMEASUREMENTCOUNT, | 2, | true, | 1, | 99, | NULL, | 0, | 0, | true, | TYPE_INT, | 3, | NULL, | (uint8_t*)"SpreadCount"},
| {(uint8_t*)"分布点间距", | 0, | COM_NUM_DEVICEPARAM_SPREADMEASUREMENTDISTANCE, | HOLDREGISTER_DEVICEPARAM_SPREADMEASUREMENTDISTANCE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"SpreadDist"},
| {(uint8_t*)"顶点距液面", | 0, | COM_NUM_DEVICEPARAM_SPREADTOPLIMIT, | HOLDREGISTER_DEVICEPARAM_SPREADTOPLIMIT, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"Top2Level"},
| {(uint8_t*)"底点距罐底", | 0, | COM_NUM_DEVICEPARAM_SPREADBOTTOMLIMIT, | HOLDREGISTER_DEVICEPARAM_SPREADBOTTOMLIMIT, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"Bot2Tank"},
| {(uint8_t*)"分布点悬停(s)", | 0, | COM_NUM_DEVICEPARAM_SPREAD_POINT_HOVER_TIME, | HOLDREGISTER_DEVICEPARAM_SPREAD_POINT_HOVER_TIME, | 2, | true, | 0, | 600, | (uint8_t*)"s", | 0, | 0, | true, | TYPE_INT, | 3, | NULL, | (uint8_t*)"PtHoverTime"},
| {(uint8_t*)"区间测量上限", | 0, | COM_NUM_DEVICEPARAM_INTERVAL_TOPLIMIT, | HOLDREGISTER_DEVICEPARAM_INTERVAL_TOPLIMIT, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"IntTop"},
| {(uint8_t*)"区间测量下限", | 0, | COM_NUM_DEVICEPARAM_INTERVAL_BOTTOMLIMIT, | HOLDREGISTER_DEVICEPARAM_INTERVAL_BOTTOMLIMIT, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"IntBottom"},
| {(uint8_t*)"保留20", | 0, | COM_NUM_DEVICEPARAM_RESERVED20, | HOLDREGISTER_DEVICEPARAM_RESERVED20, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv20"},
| {(uint8_t*)"保留21", | 0, | COM_NUM_DEVICEPARAM_RESERVED21, | HOLDREGISTER_DEVICEPARAM_RESERVED21, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv21"},
| {(uint8_t*)"密度点上限", | 0, | COM_NUM_DEVICEPARAM_WARTSILA_UPPER_DENSITY_LIMIT, | HOLDREGISTER_DEVICEPARAM_WARTSILA_UPPER_DENSITY_LIMIT, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"WUpperDen"},
| {(uint8_t*)"密度点下限", | 0, | COM_NUM_DEVICEPARAM_WARTSILA_LOWER_DENSITY_LIMIT, | HOLDREGISTER_DEVICEPARAM_WARTSILA_LOWER_DENSITY_LIMIT, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 0, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"WLowerDen"},
| {(uint8_t*)"密度点间距", | 0, | COM_NUM_DEVICEPARAM_WARTSILA_DENSITY_INTERVAL, | HOLDREGISTER_DEVICEPARAM_WARTSILA_DENSITY_INTERVAL, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 0, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"WDenStep"},
| {(uint8_t*)"最高点距液面", | 0, | COM_NUM_DEVICEPARAM_WARTSILA_MAX_HEIGHT_ABOVE_SURFACE, | HOLDREGISTER_DEVICEPARAM_WARTSILA_MAX_HEIGHT_ABOVE_SURFACE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 0, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"WMaxHeight"},
| {(uint8_t*)"保留22", | 0, | COM_NUM_DEVICEPARAM_RESERVED22, | HOLDREGISTER_DEVICEPARAM_RESERVED22, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv22"},
| {(uint8_t*)"保留23", | 0, | COM_NUM_DEVICEPARAM_RESERVED23, | HOLDREGISTER_DEVICEPARAM_RESERVED23, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv23"},
| {(uint8_t*)"保留DO高", | 0, | COM_NUM_DEVICEPARAM_ALARM_HIGH_DO, | HOLDREGISTER_DEVICEPARAM_ALARM_HIGH_DO, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"AlarmHiDO"},
| {(uint8_t*)"保留DO低", | 0, | COM_NUM_DEVICEPARAM_ALARM_LOW_DO, | HOLDREGISTER_DEVICEPARAM_ALARM_LOW_DO, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"AlarmLoDO"},
| {(uint8_t*)"保留DO三态", | 0, | COM_NUM_DEVICEPARAM_THIRD_STATE_THRESHOLD, | HOLDREGISTER_DEVICEPARAM_THIRD_STATE_THRESHOLD, | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"ThirdStateTh"},
| {(uint8_t*)"保留24", | 0, | COM_NUM_DEVICEPARAM_RESERVED24, | HOLDREGISTER_DEVICEPARAM_RESERVED24, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv24"},
| {(uint8_t*)"保留25", | 0, | COM_NUM_DEVICEPARAM_RESERVED25, | HOLDREGISTER_DEVICEPARAM_RESERVED25, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv25"},
| {(uint8_t*)"输出范围起点mA", | 0, | COM_NUM_DEVICEPARAM_CURRENT_RANGE_START_mA, | HOLDREGISTER_DEVICEPARAM_CURRENT_RANGE_START_mA, | 2, | true, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"RangeStart"},
| {(uint8_t*)"输出范围终点mA", | 0, | COM_NUM_DEVICEPARAM_CURRENT_RANGE_END_mA, | HOLDREGISTER_DEVICEPARAM_CURRENT_RANGE_END_mA, | 2, | true, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"RangeEnd"},
| {(uint8_t*)"高限报警(AO)", | 0, | COM_NUM_DEVICEPARAM_ALARM_HIGH_AO, | HOLDREGISTER_DEVICEPARAM_ALARM_HIGH_AO, | 2, | true, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"AlarmHiAO"},
| {(uint8_t*)"低限报警(AO)", | 0, | COM_NUM_DEVICEPARAM_ALARM_LOW_AO, | HOLDREGISTER_DEVICEPARAM_ALARM_LOW_AO, | 2, | true, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"AlarmLoAO"},
| {(uint8_t*)"初始电流mA", | 0, | COM_NUM_DEVICEPARAM_INITIAL_CURRENT_mA, | HOLDREGISTER_DEVICEPARAM_INITIAL_CURRENT_mA, | 2, | true, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"InitCurrent"},
| {(uint8_t*)"高位电流mA", | 0, | COM_NUM_DEVICEPARAM_AO_HIGH_CURRENT_mA, | HOLDREGISTER_DEVICEPARAM_AO_HIGH_CURRENT_mA, | 2, | true, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"HighCurrent"},
| {(uint8_t*)"低位电流mA", | 0, | COM_NUM_DEVICEPARAM_AO_LOW_CURRENT_mA, | HOLDREGISTER_DEVICEPARAM_AO_LOW_CURRENT_mA, | 2, | true, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"LowCurrent"},
| {(uint8_t*)"故障电流mA", | 0, | COM_NUM_DEVICEPARAM_FAULT_CURRENT_mA, | HOLDREGISTER_DEVICEPARAM_FAULT_CURRENT_mA, | 2, | true, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"FaultCurrent"},
| {(uint8_t*)"调试电流mA", | 0, | COM_NUM_DEVICEPARAM_DEBUG_CURRENT_mA, | HOLDREGISTER_DEVICEPARAM_DEBUG_CURRENT_mA, | 2, | true, | 0, | 3000, | (uint8_t*)"mA", | 2, | 0, | true, | TYPE_INT, | 5, | NULL, | (uint8_t*)"DebugCurrent"},
| {(uint8_t*)"保留26", | 0, | COM_NUM_DEVICEPARAM_RESERVED26, | HOLDREGISTER_DEVICEPARAM_RESERVED26, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv26"},
| {(uint8_t*)"保留27", | 0, | COM_NUM_DEVICEPARAM_RESERVED27, | HOLDREGISTER_DEVICEPARAM_RESERVED27, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv27"},
| {(uint8_t*)"标定液位值", | 0, | COM_NUM_DEVICEPARAM_CALIBRATE_OIL_LEVEL, | HOLDREGISTER_DEVICEPARAM_CALIBRATE_OIL_LEVEL, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"CalOilLvl"},
| {(uint8_t*)"标定水位值", | 0, | COM_NUM_DEVICEPARAM_CALIBRATE_WATER_LEVEL, | HOLDREGISTER_DEVICEPARAM_CALIBRATE_WATER_LEVEL, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 6, | NULL, | (uint8_t*)"CalWaterLvl"},
| {(uint8_t*)"单点测量位置", | 0, | COM_NUM_DEVICEPARAM_SP_MEAS_POSITION, | HOLDREGISTER_DEVICEPARAM_SP_MEAS_POSITION, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"SP_MeasPos"},
| {(uint8_t*)"单点监测位置", | 0, | COM_NUM_DEVICEPARAM_SP_MEAS_POSITION, | HOLDREGISTER_DEVICEPARAM_SP_MONITOR_POSITION, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"SP_MonPos"},
| {(uint8_t*)"电机运行位置", | 0, | COM_NUM_DEVICEPARAM_DENSITY_DISTRIBUTION_OIL_LEVEL, | HOLDREGISTER_DEVICEPARAM_DENSITY_DISTRIBUTION_OIL_LEVEL, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"DistOilLvl"},
| {(uint8_t*)"电机上行距离", | 0, | COM_NUM_DEVICEPARAM_MOTOR_COMMAND_DISTANCE, | HOLDREGISTER_DEVICEPARAM_MOTOR_COMMAND_DISTANCE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"MotorDist"},
| {(uint8_t*)"保留28", | 0, | COM_NUM_DEVICEPARAM_RESERVED28, | HOLDREGISTER_DEVICEPARAM_RESERVED28, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv28"},
| {(uint8_t*)"保留29", | 0, | COM_NUM_DEVICEPARAM_RESERVED29, | HOLDREGISTER_DEVICEPARAM_RESERVED29, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Rsv29"},
/* 继电器方式2报警配置（三路） */
| {(uint8_t*)"R1工作模式", | 0, | COM_NUM_DEVICEPARAM_RELAY1_OPERATING_MODE, | HOLDREGISTER_DEVICEPARAM_RELAY_OPERATING_MODE(0U), | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R1WorkMode"},
| {(uint8_t*)"R1数字源", | 0, | COM_NUM_DEVICEPARAM_RELAY1_DIGITAL_SOURCE, | HOLDREGISTER_DEVICEPARAM_RELAY_DIGITAL_SOURCE(0U), | 2, | true, | 0, | 7, | NULL, | 0, | 0, | true, | TYPE_INT, | 2, | ret_arr_word, | (uint8_t*)"R1DigSrc"},
| {(uint8_t*)"R1接点", | 0, | COM_NUM_DEVICEPARAM_RELAY1_CONTACT_TYPE, | HOLDREGISTER_DEVICEPARAM_RELAY_CONTACT_TYPE(0U), | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R1Contact"},
| {(uint8_t*)"R1报警模式", | 0, | COM_NUM_DEVICEPARAM_RELAY1_ALARM_MODE, | HOLDREGISTER_DEVICEPARAM_RELAY_ALARM_MODE(0U), | 2, | true, | 0, | 2, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R1AlarmMode"},
| {(uint8_t*)"R1故障值", | 0, | COM_NUM_DEVICEPARAM_RELAY1_ERROR_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_ERROR_VALUE(0U), | 2, | true, | 0, | 5, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R1ErrValue"},
| {(uint8_t*)"R1报警源", | 0, | COM_NUM_DEVICEPARAM_RELAY1_ALARM_SOURCE, | HOLDREGISTER_DEVICEPARAM_RELAY_ALARM_SOURCE(0U), | 2, | true, | 0, | 4, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R1AlarmSrc"},
| {(uint8_t*)"R1高高值", | 0, | COM_NUM_DEVICEPARAM_RELAY1_HH_ALARM_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_HH_ALARM_VALUE(0U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R1HHValue"},
| {(uint8_t*)"R1高值", | 0, | COM_NUM_DEVICEPARAM_RELAY1_H_ALARM_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_H_ALARM_VALUE(0U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R1HValue"},
| {(uint8_t*)"R1低值", | 0, | COM_NUM_DEVICEPARAM_RELAY1_L_ALARM_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_L_ALARM_VALUE(0U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R1LValue"},
| {(uint8_t*)"R1低低值", | 0, | COM_NUM_DEVICEPARAM_RELAY1_LL_ALARM_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_LL_ALARM_VALUE(0U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R1LLValue"},
| {(uint8_t*)"R1滞回", | 0, | COM_NUM_DEVICEPARAM_RELAY1_ALARM_HYSTERESIS, | HOLDREGISTER_DEVICEPARAM_RELAY_ALARM_HYSTERESIS(0U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R1Hysteresis"},
| {(uint8_t*)"R1阻值", | 0, | COM_NUM_DEVICEPARAM_RELAY1_DAMPING_FACTOR, | HOLDREGISTER_DEVICEPARAM_RELAY_DAMPING_FACTOR(0U), | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"R1Damping"},
| {(uint8_t*)"R1清报警", | 0, | COM_NUM_DEVICEPARAM_RELAY1_CLEAR_ALARM, | HOLDREGISTER_DEVICEPARAM_RELAY_CLEAR_ALARM(0U), | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R1Clear"},
| {(uint8_t*)"R2工作模式", | 0, | COM_NUM_DEVICEPARAM_RELAY2_OPERATING_MODE, | HOLDREGISTER_DEVICEPARAM_RELAY_OPERATING_MODE(1U), | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R2WorkMode"},
| {(uint8_t*)"R2数字源", | 0, | COM_NUM_DEVICEPARAM_RELAY2_DIGITAL_SOURCE, | HOLDREGISTER_DEVICEPARAM_RELAY_DIGITAL_SOURCE(1U), | 2, | true, | 0, | 7, | NULL, | 0, | 0, | true, | TYPE_INT, | 2, | ret_arr_word, | (uint8_t*)"R2DigSrc"},
| {(uint8_t*)"R2接点", | 0, | COM_NUM_DEVICEPARAM_RELAY2_CONTACT_TYPE, | HOLDREGISTER_DEVICEPARAM_RELAY_CONTACT_TYPE(1U), | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R2Contact"},
| {(uint8_t*)"R2报警模式", | 0, | COM_NUM_DEVICEPARAM_RELAY2_ALARM_MODE, | HOLDREGISTER_DEVICEPARAM_RELAY_ALARM_MODE(1U), | 2, | true, | 0, | 2, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R2AlarmMode"},
| {(uint8_t*)"R2故障值", | 0, | COM_NUM_DEVICEPARAM_RELAY2_ERROR_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_ERROR_VALUE(1U), | 2, | true, | 0, | 5, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R2ErrValue"},
| {(uint8_t*)"R2报警源", | 0, | COM_NUM_DEVICEPARAM_RELAY2_ALARM_SOURCE, | HOLDREGISTER_DEVICEPARAM_RELAY_ALARM_SOURCE(1U), | 2, | true, | 0, | 4, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R2AlarmSrc"},
| {(uint8_t*)"R2高高值", | 0, | COM_NUM_DEVICEPARAM_RELAY2_HH_ALARM_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_HH_ALARM_VALUE(1U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R2HHValue"},
| {(uint8_t*)"R2高值", | 0, | COM_NUM_DEVICEPARAM_RELAY2_H_ALARM_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_H_ALARM_VALUE(1U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R2HValue"},
| {(uint8_t*)"R2低值", | 0, | COM_NUM_DEVICEPARAM_RELAY2_L_ALARM_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_L_ALARM_VALUE(1U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R2LValue"},
| {(uint8_t*)"R2低低值", | 0, | COM_NUM_DEVICEPARAM_RELAY2_LL_ALARM_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_LL_ALARM_VALUE(1U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R2LLValue"},
| {(uint8_t*)"R2滞回", | 0, | COM_NUM_DEVICEPARAM_RELAY2_ALARM_HYSTERESIS, | HOLDREGISTER_DEVICEPARAM_RELAY_ALARM_HYSTERESIS(1U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R2Hysteresis"},
| {(uint8_t*)"R2阻值", | 0, | COM_NUM_DEVICEPARAM_RELAY2_DAMPING_FACTOR, | HOLDREGISTER_DEVICEPARAM_RELAY_DAMPING_FACTOR(1U), | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"R2Damping"},
| {(uint8_t*)"R2清报警", | 0, | COM_NUM_DEVICEPARAM_RELAY2_CLEAR_ALARM, | HOLDREGISTER_DEVICEPARAM_RELAY_CLEAR_ALARM(1U), | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R2Clear"},
| {(uint8_t*)"R3工作模式", | 0, | COM_NUM_DEVICEPARAM_RELAY3_OPERATING_MODE, | HOLDREGISTER_DEVICEPARAM_RELAY_OPERATING_MODE(2U), | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R3WorkMode"},
| {(uint8_t*)"R3数字源", | 0, | COM_NUM_DEVICEPARAM_RELAY3_DIGITAL_SOURCE, | HOLDREGISTER_DEVICEPARAM_RELAY_DIGITAL_SOURCE(2U), | 2, | true, | 0, | 7, | NULL, | 0, | 0, | true, | TYPE_INT, | 2, | ret_arr_word, | (uint8_t*)"R3DigSrc"},
| {(uint8_t*)"R3接点", | 0, | COM_NUM_DEVICEPARAM_RELAY3_CONTACT_TYPE, | HOLDREGISTER_DEVICEPARAM_RELAY_CONTACT_TYPE(2U), | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R3Contact"},
| {(uint8_t*)"R3报警模式", | 0, | COM_NUM_DEVICEPARAM_RELAY3_ALARM_MODE, | HOLDREGISTER_DEVICEPARAM_RELAY_ALARM_MODE(2U), | 2, | true, | 0, | 2, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R3AlarmMode"},
| {(uint8_t*)"R3故障值", | 0, | COM_NUM_DEVICEPARAM_RELAY3_ERROR_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_ERROR_VALUE(2U), | 2, | true, | 0, | 5, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R3ErrValue"},
| {(uint8_t*)"R3报警源", | 0, | COM_NUM_DEVICEPARAM_RELAY3_ALARM_SOURCE, | HOLDREGISTER_DEVICEPARAM_RELAY_ALARM_SOURCE(2U), | 2, | true, | 0, | 4, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R3AlarmSrc"},
| {(uint8_t*)"R3高高值", | 0, | COM_NUM_DEVICEPARAM_RELAY3_HH_ALARM_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_HH_ALARM_VALUE(2U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R3HHValue"},
| {(uint8_t*)"R3高值", | 0, | COM_NUM_DEVICEPARAM_RELAY3_H_ALARM_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_H_ALARM_VALUE(2U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R3HValue"},
| {(uint8_t*)"R3低值", | 0, | COM_NUM_DEVICEPARAM_RELAY3_L_ALARM_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_L_ALARM_VALUE(2U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R3LValue"},
| {(uint8_t*)"R3低低值", | 0, | COM_NUM_DEVICEPARAM_RELAY3_LL_ALARM_VALUE, | HOLDREGISTER_DEVICEPARAM_RELAY_LL_ALARM_VALUE(2U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R3LLValue"},
| {(uint8_t*)"R3滞回", | 0, | COM_NUM_DEVICEPARAM_RELAY3_ALARM_HYSTERESIS, | HOLDREGISTER_DEVICEPARAM_RELAY_ALARM_HYSTERESIS(2U), | 2, | false, | 0, | 0, | NULL, | 1, | 0, | true, | TYPE_FLOAT, | 8, | NULL, | (uint8_t*)"R3Hysteresis"},
| {(uint8_t*)"R3阻值", | 0, | COM_NUM_DEVICEPARAM_RELAY3_DAMPING_FACTOR, | HOLDREGISTER_DEVICEPARAM_RELAY_DAMPING_FACTOR(2U), | 2, | false, | 0, | 0, | NULL, | 0, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"R3Damping"},
| {(uint8_t*)"R3清报警", | 0, | COM_NUM_DEVICEPARAM_RELAY3_CLEAR_ALARM, | HOLDREGISTER_DEVICEPARAM_RELAY_CLEAR_ALARM(2U), | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"R3Clear"},
| {(uint8_t*)"参数版本号", | 0, | COM_NUM_DEVICEPARAM_PARAM_VERSION, | HOLDREGISTER_DEVICEPARAM_PARAM_VERSION, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 4, | NULL, | (uint8_t*)"ParamVer"},
| {(uint8_t*)"结构体大小", | 0, | COM_NUM_DEVICEPARAM_STRUCT_SIZE, | HOLDREGISTER_DEVICEPARAM_STRUCT_SIZE, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 4, | NULL, | (uint8_t*)"StructSize"},
| {(uint8_t*)"魔术字", | 0, | COM_NUM_DEVICEPARAM_MAGIC, | HOLDREGISTER_DEVICEPARAM_MAGIC, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"Magic"},
| {(uint8_t*)"参数CRC32", | 0, | COM_NUM_DEVICEPARAM_CRC, | HOLDREGISTER_DEVICEPARAM_CRC, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 8, | NULL, | (uint8_t*)"ParamCRC"},
/* ==================== 下发指令参数（保持你原来的寄存器宏）==================== */
| {(uint8_t*)"标定液位值", | 0, | COM_NUM_CAL_OIL, | HOLDREGISTER_DEVICEPARAM_CALIBRATE_OIL_LEVEL, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"CalOilLvl"},
| {(uint8_t*)"修正液位值", | 0, | COM_NUM_CORRECTION_OIL, | HOLDREGISTER_DEVICEPARAM_CALIBRATE_OIL_LEVEL, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"CalOilLvl"},
| {(uint8_t*)"标定水位值", | 0, | COM_NUM_CALIBRATE_WATER, | HOLDREGISTER_DEVICEPARAM_CALIBRATE_WATER_LEVEL, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"CalWaterLvl"},
| {(uint8_t*)"单点测量位置", | 0, | COM_NUM_SINGLE_POINT, | HOLDREGISTER_DEVICEPARAM_SP_MEAS_POSITION, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"SP_MeasPos"},
| {(uint8_t*)"单点监测位置", | 0, | COM_NUM_SP_TEST, | HOLDREGISTER_DEVICEPARAM_SP_MONITOR_POSITION, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"SP_MonPos"},
| {(uint8_t*)"电机运行位置", | 0, | COM_NUM_RUN_TO_POSITION, | HOLDREGISTER_DEVICEPARAM_DENSITY_DISTRIBUTION_OIL_LEVEL, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"DistOilLvl"},
| {(uint8_t*)"电机上行距离", | 0, | COM_NUM_RUNUP, | HOLDREGISTER_DEVICEPARAM_MOTOR_COMMAND_DISTANCE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"MotorDist"},
| {(uint8_t*)"电机下行距离", | 0, | COM_NUM_RUNDOWN, | HOLDREGISTER_DEVICEPARAM_MOTOR_COMMAND_DISTANCE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"MotorDist"},
| {(uint8_t*)"电机上行距离", | 0, | COM_NUM_FORCE_RUNUP, | HOLDREGISTER_DEVICEPARAM_MOTOR_COMMAND_DISTANCE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"MotorDist"},
| {(uint8_t*)"电机下行距离", | 0, | COM_NUM_FORCE_RUNDOWN, | HOLDREGISTER_DEVICEPARAM_MOTOR_COMMAND_DISTANCE, | 2, | false, | 0, | 0, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"MotorDist"},
/* ==================== CPU3 本机参数（保持你原来的寄存器宏）==================== */
| {(uint8_t*)"屏幕程序版本", | 0, | COM_NUM_PARA_LOCAL_LEDVERSION, | HOLDREGISTER_CPU3_LED_VERSION, | 2, | false, | 0, | 0, | NULL, | 0, | 0, | false, | TYPE_INT, | 4, | NULL, | (uint8_t*)"LedVer"},
| {(uint8_t*)"语言", | 0, | COM_NUM_PARA_LANG, | HOLDREGISTER_CPU3_LANGUAGE, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"Lang"},
| {(uint8_t*)"液位数据源", | 0, | COM_NUM_SCREEN_SOURCE_OIL, | HOLDREGISTER_CPU3_SRC_OIL, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"SrcOil"},
| {(uint8_t*)"水位数据源", | 0, | COM_NUM_SCREEN_SOURCE_WATER, | HOLDREGISTER_CPU3_SRC_WATER, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"SrcWater"},
| {(uint8_t*)"密度数据源", | 0, | COM_NUM_SCREEN_SOURCE_D, | HOLDREGISTER_CPU3_SRC_D, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"SrcD"},
| {(uint8_t*)"温度数据源", | 0, | COM_NUM_SCREEN_SOURCE_T, | HOLDREGISTER_CPU3_SRC_T, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"SrcT"},
| {(uint8_t*)"液位手输值", | 0, | COM_NUM_SCREEN_INPUT_OIL, | HOLDREGISTER_CPU3_IN_OIL, | 2, | true, | -999999, | 999999, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"InOil"},
| {(uint8_t*)"水位手输值", | 0, | COM_NUM_SCREEN_INPUT_WATER, | HOLDREGISTER_CPU3_IN_WATER, | 2, | true, | -999999, | 999999, | (uint8_t*)"mm", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"InWater"},
| {(uint8_t*)"密度手输值", | 0, | COM_NUM_SCREEN_INPUT_D, | HOLDREGISTER_CPU3_IN_D, | 2, | true, | 0, | 2000, | (uint8_t*)"kg/m3", | 3, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"InD"},
| {(uint8_t*)"密度手输上传", | 0, | COM_NUM_SCREEN_INPUT_D_SWITCH, | HOLDREGISTER_CPU3_IN_D_SW, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"InDSw"},
| {(uint8_t*)"温度手输值", | 0, | COM_NUM_SCREEN_INPUT_T, | HOLDREGISTER_CPU3_IN_T, | 2, | true, | -500, | 2000, | (uint8_t*)"℃", | 1, | 0, | true, | TYPE_INT, | 7, | NULL, | (uint8_t*)"InT"},
| {(uint8_t*)"小数点位数", | 0, | COM_NUM_SCREEN_DECIMAL, | HOLDREGISTER_CPU3_DECIMAL, | 2, | true, | 0, | 4, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"Decimal"},
| {(uint8_t*)"屏幕密码", | 0, | COM_NUM_SCREEN_PASSWARD, | HOLDREGISTER_CPU3_PASSWORD, | 2, | true, | 0, | 9999, | NULL, | 0, | 0, | true, | TYPE_INT, | 4, | NULL, | (uint8_t*)"ScrPwd"},
| {(uint8_t*)"是否息屏", | 0, | COM_NUM_SCREEN_OFF, | HOLDREGISTER_CPU3_OFF_TIME, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | NULL, | (uint8_t*)"ScreenOff"},
| {(uint8_t*)"COM1波特率", | 0, | COM_NUM_CPU3_COM1_BAUDRATE, | HOLDREGISTER_CPU3_COM1_BAUD, | 2, | true, | 0, | 5, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C1Baud"},
| {(uint8_t*)"COM1数据位", | 0, | COM_NUM_CPU3_COM1_DATABITS, | HOLDREGISTER_CPU3_COM1_DATABITS, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C1Data"},
| {(uint8_t*)"COM1校验", | 0, | COM_NUM_CPU3_COM1_PARITY, | HOLDREGISTER_CPU3_COM1_PARITY, | 2, | true, | 0, | 2, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C1Parity"},
| {(uint8_t*)"COM1停止位", | 0, | COM_NUM_CPU3_COM1_STOPBITS, | HOLDREGISTER_CPU3_COM1_STOPBITS, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C1Stop"},
| {(uint8_t*)"COM1协议", | 0, | COM_NUM_CPU3_COM1_PROTOCOL, | HOLDREGISTER_CPU3_COM1_PROTO, | 2, | true, | 0, | 4, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C1Proto"},
| {(uint8_t*)"COM2波特率", | 0, | COM_NUM_CPU3_COM2_BAUDRATE, | HOLDREGISTER_CPU3_COM2_BAUD, | 2, | true, | 0, | 5, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C2Baud"},
| {(uint8_t*)"COM2数据位", | 0, | COM_NUM_CPU3_COM2_DATABITS, | HOLDREGISTER_CPU3_COM2_DATABITS, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C2Data"},
| {(uint8_t*)"COM2校验", | 0, | COM_NUM_CPU3_COM2_PARITY, | HOLDREGISTER_CPU3_COM2_PARITY, | 2, | true, | 0, | 2, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C2Parity"},
| {(uint8_t*)"COM2停止位", | 0, | COM_NUM_CPU3_COM2_STOPBITS, | HOLDREGISTER_CPU3_COM2_STOPBITS, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C2Stop"},
| {(uint8_t*)"COM2协议", | 0, | COM_NUM_CPU3_COM2_PROTOCOL, | HOLDREGISTER_CPU3_COM2_PROTO, | 2, | true, | 0, | 4, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C2Proto"},
| {(uint8_t*)"COM3波特率", | 0, | COM_NUM_CPU3_COM3_BAUDRATE, | HOLDREGISTER_CPU3_COM3_BAUD, | 2, | true, | 0, | 5, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C3Baud"},
| {(uint8_t*)"COM3数据位", | 0, | COM_NUM_CPU3_COM3_DATABITS, | HOLDREGISTER_CPU3_COM3_DATABITS, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C3Data"},
| {(uint8_t*)"COM3校验", | 0, | COM_NUM_CPU3_COM3_PARITY, | HOLDREGISTER_CPU3_COM3_PARITY, | 2, | true, | 0, | 2, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C3Parity"},
| {(uint8_t*)"COM3停止位", | 0, | COM_NUM_CPU3_COM3_STOPBITS, | HOLDREGISTER_CPU3_COM3_STOPBITS, | 2, | true, | 0, | 1, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C3Stop"},
| {(uint8_t*)"COM3协议", | 0, | COM_NUM_CPU3_COM3_PROTOCOL, | HOLDREGISTER_CPU3_COM3_PROTO, | 2, | true, | 0, | 4, | NULL, | 0, | 0, | true, | TYPE_INT, | 1, | ret_arr_word, | (uint8_t*)"C3Proto"},
};